INSTRUCTIONS

This is a safe work procedure template for:

- Fully enclosed or partially enclosed x-ray analysis equipment

You can edit any of the red text in this template so that it reflects your own safe work practices and procedures (change back to black text before submitting). The remainder of the document should not be changed without consulting with your local Radiation Safety Supervisor or the University Radiation Safety Officer.

Delete this instruction page before submitting this SWP with your radiation project application.

| Safe Work Procedure |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Scope |  |  |  |  |  |  |  |
| This document defines the Safe Work Procedure for [Insert procedure name]. |  |  |  |  |  |  |  |
| Application |  |  |  |  |  |  |  |
| XXX |  |  |  |  |  |  |  |
| Safety Hazards and Risk Controls |  |  |  |  |  |  |  |
| Hazard |  |  |  |  | Control Strategy |  |  |
| - Damage to the skin by making physical contact with the beam. |  |  |  |  | Ensure all protective guards are in place and that any electronic safety measures (e.g. interlocks) are operational. Maintain safe distance from any partially enclosed x-ray beam operating. |  |  |
| - [Insert] |  |  |  |  | [Insert] |  |  |
| - [Insert] |  |  |  |  | [Insert] |  |  |
| - [Insert] |  |  |  |  | [Insert] |  |  |
| Personal Protective Equipment (PPE) |  |  |  |  |  |  |  |
| Appropriate personal protective equipment must be provided and used. (Delete PPE not required). PPE must be appropriately selected, individually fitted and workers trained in correct use and maintenance: |  |  |  |  |  |  |  |
| Footwear |  | Safety footwear must be worn at all times. | Safety Glasses |  |  |  | Laser safety glasses, goggles &/or face visor must be worn at all times. |
| Protective Gloves |  | Safety gloves must be worn at all times. | Protective Clothing |  |  |  | Close fitting/protective clothing must be worn. |
| Helmet |  | Helmets must be worn on construction sites at all times. | Hearing Protection |  |  |  | Hearing protection may be required for some operations. |
| Respiratory Protection |  | Respiratory protection devices may be required for some operations. | Hi Visibility Vests |  |  |  | Vests must be worn at all times |
| Hair |  | Long and loose hair must be contained. |  |  |  |  |  |
| Instrument Calibration |  |  |  |  |  |  |  |
| Instrument Name |  |  |  | Date of Calibration |  |  |  |
| [Insert] |  |  |  | [Insert] |  |  |  |
| [Insert] |  |  |  | [Insert] |  |  |  |
| [Insert] |  |  |  | [Insert] |  |  |  |
| Roles and Responsibilities |  |  |  |  |  |  |  |
| Title | Role |  |  |  |  |  |  |
| - Radiation Safety Supervisor (or Head of School if RSS role vacant) | The Radiation Safety Supervisor (RSS) is an administrative role appointed by the Head of School to act on their behalf to ensure radiation projects within the area follow University requirements and processes. Responsible for authorising entry to the radiation laboratory, and maintaining a list of those authorised personnel. Ensure that personnel are aware of their responsibilities in relation to this SWP. |  |  |  |  |  |  |
| - Licensee | The licensee has a licence granted by the WA State Government to work and supervise others on projects related to the conditions under their licence. The licensee has legal responsibilities for the day-to-day radiation safety aspects of the work carried out under their licence. If personnel without appropriate training require access to the radiation area (e.g. laboratory managers, technicians, students etc.) the licensee must initially ensure the area is safe before allowing access and then supervise those personnel at all times. |  |  |  |  |  |  |
| - Chief Investigator | The Chief Investigator (CI) of a project using radiation equipment/substances in the radiation laboratory must ensure the project is approved by the University Radiation Safety Officer or Committee prior to commencing work. If approval is granted the CI must ensure work is conducted in line with this SWP and any conditions under the project approval. The CI is responsible for ensuring a risk assessment has been undertaken for the work related to their project. |  |  |  |  |  |  |
| - Radiation workers (staff and students using radiation) | Radiation workers must have undertaken relevant radiation safety training and either work under the general supervision of a licensee (i.e. the licensee is on site but not necessarily present in the room during operation) or must have their own laser licence issued by the WA State Government. Conduct work in accordance with this SWP and the conditions of approval for the project they are working on. When leaving the laboratory, ensure the radiation laboratory is safe for other personnel to enter. If experiments are ongoing, ensure adequate signage is present to warn other personnel. Immediately report any incidents or hazards. |  |  |  |  |  |  |
| - Other personnel | Other personnel that require access to the radiation laboratory (e.g. laboratory managers, technicians, students etc.) can only enter the area with the authorisation of the RSS and under the immediate supervision of a licensee. |  |  |  |  |  |  |
| Procedure |  |  |  |  |  |  |  |
| Pre-operational Safety Check & General Working Rules |  |  |  |  |  |  |  |
| - The unit must not be operated if any of the warning lights are not operational. |  |  |  |  |  |  |  |
| - No x-ray tube may be energised; while outside its protective tube housing or, with an unshielded aperture in the tube head or protective barrier. |  |  |  |  |  |  |  |
| - No sample, collimator or analysing crystal shall be changed or adjusted whilst a primary x-ray beam passes through that collimator or is incident on that sample or crystal unless; the sample, collimator or crystal, during and after the change or adjustment is within a shielded enclosure and the change or adjustment is performed by remote means from outside the enclosure. |  |  |  |  |  |  |  |
| - Visual alignments or adjustments must not be carried out whilst the x-ray tube is energised, unless a viewing system is used which is shielded or designed to prevent exposure of the eye or other parts of the body to the primary beam. |  |  |  |  |  |  |  |
| - XRA equipment is not to be operated with interlocks inactivated without prior approval from the statutory authority, the WA Radiological Council. |  |  |  |  |  |  |  |
| - Service may only be carried out by a service technician with a suitable Radiological Council licence. |  |  |  |  |  |  |  |
| - Radiation leakage surveys are required - Fully enclosed units quarterly, partly enclosed units monthly. Records must be maintained. |  |  |  |  |  |  |  |
| - Accidents and incidents must be reported to Health, Safety and Emergency Management via the CHARM system |  |  |  |  |  |  |  |
| Operational Safety Check |  |  |  |  |  |  |  |
| - Inspect for any potentially hazardous situations arising from x-ray beams that may be emitted due to equipment defect, misalignment or any other reason. |  |  |  |  |  |  |  |
| - Conduct a brief hazard identification for the laboratory as a whole. |  |  |  |  |  |  |  |
| - Use the required PPE. |  |  |  |  |  |  |  |
| Startup and Operating Instructions |  |  |  |  |  |  |  |
| - Exposure to the primary x-ray beam must be avoided at all times. |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| Housekeeping |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| - [Insert] |  |  |  |  |  |  |  |
| Reference Material and Attachments |  |  |  |  |  |  |  |
| - Code of Practice for Protection Against Ionizing Radiation Emitted from X-ray Analysis Equipment, NHMRC 1984 - Manufacturer’s standard operating procedure for the x-ray equipment. |  |  |  |  |  |  |  |
| Authorisation |  |  |  |  |  |  |  |
| Approved for use by Curtin University Radiation Safety Officer/Committee: Approval number RSC….-.. |  |  |  |  |  |  |  |